# CA6240 firmware update trough a PC

### Requirements

  - IBM PC or compatible (Pentium or better)

  - Microsoft(R) Windows(R) 98 / NT 4 / 2000 / Me / XP / Vista

  - USB/RS-232 optical link

  - FLASHSTA.EXE  program

The program “Flash Programmer M16C Flash Starter” with FLASHSTA.EXE as executable program can be downloaded with following address Internet: http://www.renesas.com/fmwk.jsp?cnt=/download_search_results.jsp&fp=/support/downloads/download_results&layerId=1216

  - The two following files shall be in the same directory than the program Flashsta.exe

- the file of the loaded firmware (filename with *.mot extension)

- the file with *.id extension.

### Prepare to download

Before downloading the firmware , the user shall verify that the battery is full ( the charge is made while the instrument is OFF)

The user shall make a reset and then put the instrument in a particular state for downloading as described below :

  - To Make a reset of the instrument

  - The rotary commutator must be on the position ‘4000m/100mA’

  - Press at the same time on the buttons ‘CLR’ and ‘MR’ and ‘DISPLAY’ and then unpress.

  - Prepare to download

- The rotary commutator must be on the position ‘4000m/100mA’

- Press at the same time on the buttons ‘CLR’ and ‘MEM’ and ‘DISPLAY’ and then unpress.

### Downloading the firmware into the instrument

### Launching of the downloading program :

### The downloading program is self-contained and does’nt required any installation.

### Simply open FLASHTA.EXE.

### A first window appears

- In Select Program  «  Internal Flash memory » is checked by default

- Choose the COM port to be used.

- Click on OK

### A second window appears.

- Click on the button Refer…

- Enter the complete path of the file with filename extension « .mot » .

- When the filename is specified, the ID boxes will be fulfilled automatically.

- Click on OK.

### A third window appears.

Erase by Click on the button Erase.

Click on OK.

Then click on the button Program…

For downloading, enter the start and stop addresses according to the procedure below and then press the button Ok (2 times):

### When the downloading is finished, make a reset of the instrument

### as described below :

  - The rotary commutator must be on the position ‘4000m/100mA’

  - Press at the same time on the buttons ‘CLR’ and ‘MR’ and ‘DISPLAY’ and then unpress.